TEST SUMMARY REPORT: VITALFLOW SYSTEM

IEEE 829-1998 Standard Format

1. Test Summary Report Identifier TSR-VF-001

(Refers to Software Level: VitalFlow v1.0)

2. Summary

This report summarizes the testing activities for VitalFlow v1.0. It confirms that all items defined in Test Plan TP-VF-001 were tested. The specific version tested was Build 1.0.2.

- Environment: Windows 11, WSL, Spin 6.5.1.

3. Variances

- Deviation: The "Predictive Trend Arrow" feature was excluded from testing.

- Reason: Development team did not finish the algorithm in time. A Change Request (CR-101) was filed to move this to Release 2.0.

4. Comprehensiveness Assessment

Testing is considered effective and comprehensive regarding safety objectives.

- Coverage: 100% of Safety Objectives were verified using Promela.

- Objectives Covered: All High and Medium priority features were executed.

- Omitted Objectives: "Predictive Trend" (as noted in Variances).

5. Summary of Results

The overall status of incidents highlights a stable backend but an unstable UI.

5.1. Design Verification Log (Spin Tool Output)

The following log confirms that the Safety and Liveness properties were verified.

SCREENSHOT

Result: PASS

5.2. Incident Summary

- Total Incidents: 4.

- By Severity: 1 Critical (Fixed), 1 High (Open), 1 Medium (Fixed), 1 Low (Closed).

- Defect Patterns: Most defects were found in the Bluetooth connectivity module.

6. Evaluation

Based on the testing results, the software quality is rated as High Safety / Medium Usability.

- Failure Likelihood: Low for critical safety features (verified by Formal Methods).

- Limitations: The "Bluetooth Disconnection" issue (INC-01) remains open.

- Good News: The core pumping logic had zero defects during regression.

- Recommendation: The software is acceptable for Beta Testing under clinical supervision, but general release should be held until INC-01 is resolved.

7. Summary of Activities

- Staff Time: Actual effort for Promela modeling exceeded planned effort by 20% due to tool complexity.

- Costs: Remained within budget; no overtime required.

- Test Environment: Available 100% of the time; no delays caused by hardware.