TEST SUMMARY REPORT: VITALFLOW JAVA SYSTEM

IEEE 829-1998 Standard Format

1. Test Summary Report Identifier

ID: VF-JAVA-TSR-2025-001

Level: System Implementation Summary

Related Software: VitalFlow Java Application v1.0

This document serves as the final Test Summary Report for the verification of the VitalFlow Insulin Pump Java implementation. It correlates directly with Test Plan VF-JAVA-TP-2025-001 and summarizes the results of the JUnit-style testing and manual GUI verification activities performed in the current cycle.

The report covers the validation of the safety logic in Controller.java and technical implementation of the BlockingQueue communication model. It serves as the formal “Certificate of Correctness” for the Java code before final deployment.

2. Summary

The software verified in this cycle is the VitalFlow Java Application v1.0. The primary test items included the concurrent interactions between the Sensor, Controller and Pump threads. We specifically validated the system against the critical safety requirement: ensuring that insulin is never delivered when glucose levels are below 70 mg/dL.

The test environment utilized JDK 23 on a Windows laptop. Automated tests were executed using the SimpleTestRunner utility, which scanned the com.vitalflow.test package for methods annotated with @Test.

This summary confirms that the scope of testing was adhered to, covering the “Critical High” features of Hypoglycemia Protection and the “Medium” risk features of Hyperglycemia Response. The “Reservoir Depletion Alarm” logic was also included in the final verification pass.

3. Variances

There was one significant deviation from the original Test Plan regarding the use of external testing frameworks. Originally, the plan called for JUnit 5. However, we utilized a custom SimpleTestRunner class which replicates the core functionality of JUnit. This variance did not impact the quality of testing as Assertions (assertTrue, assertFalse) were manually verified against the expected logic.

A second variance occurred regarding GUI automation. While the plan mentioned TestFX, we pivoted to manual exploratory testing for the GUI to prioritize fixing race conditions in the  shared object found during early integration runs.

4. Comprehensiveness Assessment

The testing process is assessed as comprehensive regarding functional logic. Every logic branch in Controller.decidePumpAction() was exercised through the controllerLogicTest suite, covering boundaries at 69, 70, 199, and 200 mg/dL.

Regarding concurrency, we achieved high coverage by running the application for extended periods (over 5 minutes) under “Stress” conditions (1-second sensor updates). This allowed us to observe several thousand message-passing cycles between threads. While we cannot mathematically guarantee the absence of all race conditions (unlike the previous Promela model), the empirical stability observed during these runs provides high confidence in the implementation.

5. Summary of Results

The verification cycle identified a total of 2 defects:

- Defect #01 (Critical): A InterruptedException in the Pump thread caused the insulin delivery loop to crash when the GUI window was moved rapidly. Status: FIXED.

- Defect #02 (High): The insulinReservoir value in the GUI did not update in real-time due to a missing SystemState refresh in the Controller loop. Status: FIXED.

The final test suite run produced:

- Total Tests: 7

- Passed: 7

- Failed: 0

- Errors: 0

6. Evaluation

Based on the zero-error result of the final automated test run and the stability of the long-duration integration test, the VitalFlow Java code is evaluated as High Quality and suitable for demonstration. The core safety invariant (No Pumping < 70) holds true in all test cases.

The “Failure Likelihood” regarding logic errors is assessed as low. However, we note that the system relies on the JVM’s thread scheduler. In extremely resource-constrained environments, thread starvation could theoretically delay an emergency shutoff by a few milliseconds.

7. Summary of Activities

The testing activities consumed a total of 12 hours.

- 4 hours: Developing and refining the SimpleTestRunner and ControllerLogicTest.

- 6 hours: Integration testing and manual GUI verification.

- 2 hours: Defect resolution and final reporting.

The project scope remained stable, allowing to focus on the transition from the formal model to the physical Java implementation.